Investigation of Two Models to Set and Evaluate Quality Targets for HbA1c: Biological Variation and Sigma-metrics

Authors:

Cas Weykamp (1,2) (*)

Garry John (3)

Philippe Gillery (4)

Emma English (5)

Linong Ji (6)

Erna Lenters-Westra (7,8)

Randie R. Little (9)

Gojka Roglic (10)

David B. Sacks (11)

Izumi Takei (12)

On behalf of the IFCC Task Force on Implementation of HbA1c Standardization (TF-HbA1c)

- Department of Clinical Chemistry, Queen Beatrix Hospital, Winterswijk, the Netherlands

- European Reference Laboratory, Location Queen Beatrix Hospital, Winterswijk, the Netherlands

- Norfolk and Norwich University Hospital, Norwich, UK

- Laboratory of Pediatric Biology and Research, University Hospital of Reims, France

- School of Medicine, University of Nottingham, Royal Derby Hospital Site, UK

- Peking University People’s Hospital, Beijing, China

- Department of Clinical Chemistry, Isala Clinics, Zwolle, the Netherlands

- European Reference Laboratory, Location Isala, Zwolle, the Netherlands

- Departments of Pathology and Child Health, University of Missouri School of Medicine, Columbia, USA

- Department of Management of Noncommunicable Diseases, World Health Organization, Geneva, Switzerland

- Department of Laboratory Medicine, National Institutes of Health, Bethesda, USA

- Diabetes and Endocrine Department, Ichikawa General Hospital, Ichikawa, Japan

* Address correspondence to this author at:

Queen Beatrix Hospital, Department of Clinical Chemistry, Beatrixpark 1, 7101 BN Winterswijk, the Netherlands. Tel +31 543 544774; Fax +31 543 524265; e-mail c.w.weykamp@skbwinterswijk.nl

Running Head: Quality Targets for HbA1c

Keywords: HbA1c, Quality Targets, Biological Variation, Sigma-metrics, Diabetes

Abbreviations: IFCC, International Federation of Clinical Chemistry and Laboratory Medicine; TF- HbA1c, Task Force HbA1c; TAE, total allowable error; TE, total error; SM, Sigma-metrics model; BV, biological variation model; EQA, external quality assessment; PT, proficiency testing; POCT, Point of Care Testing;  SI, Système International; IDF, International Diabetes Federation; EASD, European Association for the Study of Diabetes; ADA, American Diabetes Association; NGSP, National Glycohemoglobin Standardization Program; CAP, College of American Pathologists; σ, sigma; CVi, intra-individual biological variation; CVg, inter-individual biological variation; B, bias; I, imprecision.

Abstract

BACKGROUND: A major objective of the IFCC Task Force on implementation of HbA1c standardization is to develop a model to define quality targets for HbA1c.

METHODS: Two generic models, the Biological Variation and Sigma-metrics model, are investigated. Variables in the models are filled for HbA1c and data of EQA/PT programs are used to evaluate the suitability of the models to set and evaluate quality targets within and between laboratories.

RESULTS: In the biological variation model 48% of individual laboratories and none of the 26 instrument groups met the minimum performance criterion. In the Sigma-metrics model, with a total allowable error (TAE) set at 5 mmol/mol (0.46% NGSP) 77% of the individual laboratories and 12 of 26 instrument groups met the 2 sigma criterion.

CONCLUSION: The Biological Variation and Sigma-metrics model were demonstrated to be suitable for setting and evaluating quality targets within and between laboratories. The Sigma-metrics model is more flexible as both the TAE and the risk of failure can be adjusted to requirements related to e.g. use for diagnosis/monitoring or requirements of (inter)national authorities. With the aim of reaching international consensus on advice on quality targets for HbA1c, the Task Force suggests the Sigma-metrics model as the model of choice with default values of 5 mmol/mol (0.46%) for TAE, and risk levels of 2 and 4 sigma for routine laboratories and laboratories performing clinical trials, respectively. These goals should serve as a starting point for discussion with international stakeholders in the field of diabetes.

Introduction

A major objective of the International Federation of Clinical Chemistry and Laboratory Medicine (IFCC) Task Force on Implementation of HbA1c Standardization (TF- HbA1c) is to “Develop quality targets for the measurement of HbA1c,  and on the basis of these targets, and in conjunction with professional bodies, advise on the use of HbA1c for monitoring, diagnosis and screening of diabetes and glucose intolerance” (1). This paper addresses the development of a model to set and evaluate quality targets. A suitable model should be applicable both within and between laboratories. The IFCC is an international organization that advises on quality targets and therefore the model should be generic and give local regulatory bodies the option to fill specific criteria related to e.g. financial and technical resources.

Table 1 summarizes the key elements considered in the development of the models.  A quality target consists of a value assigned by an approved reference measurement procedure, with tolerance limits derived from the quality concept of Total Allowable Error (TAE) (2). In this concept Total Error (TE) is derived from bias and imprecision and compared with the TAE to reveal a pass or fail. The suitability of two TAE based analytical-mathematical models is investigated here: the “Biological Variation” model (BV) (3,4) and the “Sigma-metrics” model (SM) (5). To set quality targets both models require alignment with the Stockholm Hierarchy for analytical quality (6) to define TAE. In this hierarchy there are five levels: a) clinical outcomes in specific clinical settings, b) data derived from biological variation, c) published professional recommendations, d) performance goals set by regulatory bodies or organizers of External Quality Assessment (EQA) schemes, and e) current state of the art. For the SM model all levels can be used, whereas only level b applies to the BV model. Both models have a grading system to categorize the risk of not achieving the set TAE. In the SM model the risk is expressed in sigma (σ) units (the lower σ, the higher the risk of not achieving the set TAE) (5). The BV model has three categories, namely the optimum, desirable and minimum level of performance (3,4).

Several considerations may have an impact on the specifically defined criteria within both generic models: the aim of the test (e.g. diagnosis versus monitoring), test environment (e.g. Central Lab versus POCT), technical limitations (what is the state of the art), economic limitations (what cost is economically affordable), practical considerations (workload and availability of  bias and imprecision data), and clinical needs (4).

This study evaluates the suitability of the BV and SM models within and between laboratories by using data from External Quality Assessment/Proficiency Testing (EQA/PT) programs in the US and Europe. In compliance with the Consensus Statement on HbA1c of the International Diabetes Federation (IDF), European Association for the Study of Diabetes (EASD), American Diabetes Association (ADA) and IFCC, data will be expressed in Système International (SI) units with National Glycohemoglobin Standardization Program (NGSP) units (7) in brackets. It should be noted that the mathematical relationships between the IFCC and NGSP systems are complicated; conversion factors are summarized in table 2 and details can be found in reference (8).

Methods

Biological Variation Model

The outcome of the BV model varies with HbA1c concentration (8) and necessitates the choice of one HbA1c concentration for evaluation. 50 mmol/mol (6.7%) is selected as most relevant as this concentration is close to and in-between the clinical decision limits for diagnosis of diabetes (48 mmol/mol; 6.5%) and a common target for treatment (53 mmol/mol; 7.0%) (9). TE and TAE are calculated with equations 1 and  2 in Table 2. Intra- and inter-individual biological variation is taken from the table published by Ricos (10). TAE is expressed in relative and absolute units .  Fig. 1A is the graphical representation of the limits for minimum, desirable and optimum performance. There are two extreme options to achieve an adequate performance level: a perfect bias with a relatively poor imprecision or a perfect imprecision with a relatively poor bias. As an example, the performance of a laboratory with a bias of 1.0 mmol/mol (0.09%) and an imprecision of 1.5% (1.0%) resulting in a TE of 2.5 mmol/mol (0.23%) is drawn in the figure. This lab meets the BV criterion for minimum performance but fails the desirable and optimum criteria.

Sigma-metrics Model

Sigma-metrics provides a universal benchmark for process performances. The performance of any process can be characterized on a “Sigma scale”. Values range from 2 to 6. A higher (4-6) sigma level implies a very low risk to fail (thus a very reliable process) but also a very narrow target range (thus difficult to achieve). A lower (2-4) level means an easier to achieve performance level but also a higher risk to fail (i.e. more results ouside the set TAE value). The target sigma level set should balance optimum quality and risk to fail (5). In the SM model TAE is set according to one of the levels of the Stockholm Hierarchy Consensus (6) along with the risk of failing this goal. The performance of a laboratory is expressed in a sigma score and is calculated with equation 3 in Table 2. This same equation can be used to calculate allowable bias and imprecision at a defined sigma level. This is graphically represented in Fig.1B. To evaluate the model a TAE of 5 mmol/mol (0.46%) is chosen. This is based on the difference in HbA1c result in two consecutive HbA1c tests that guide clinicians to change therapy (11) and is therefore a clinical decision limit (Stockholm Hierarchy level 1). Five mmol/mol (0.46%) is also approximately the difference between the upper level of persons at low risk for diabetes and the decision limit for diagnosis of diabetes (48 mmol/mol; 6.5%) (9,12). As with the BV model there are extremes of bias and imprecision that allow compliance with the TAE: a high bias with a low imprecision or vice versa. Fig.1B also shows the performance of the same laboratory as in the BV model: the TE of 2.5 mmol/mol (0.17%) results in a score of 5.3 sigma.

Comparison Models

In Fig. 1C both models are combined. It is evident that BV limits are tighter than SM limits. This is not surprising as the arbitrary set TAE is higher in the SM model (5 mmol/mol; 0.46%) than those in the BV model (3.4/2.2/1.1 mmol/mol; 0.31, 0.20, 0.10%). In the SM model there is free choice to define TAE and sigma level. The impact is demonstrated in Fig.1D. When a TAE of 3 mmol/mol (0.27%)  is set with an acceptable risk of 4σ, the lab fails. But with a TAE of 5 mmol/mol (0.46%) and an acceptable risk of 2σ, the lab passes. In the BV model both TAE and risk categories are fixed by intra- and inter-individual variation. The values from the Ricos tables may be revised following recent literature (13,14,15) which will of course lead to different (tighter) quality targets.

Results

Evaluation between laboratories

Data from the College of American Pathologists (CAP) Survey 2014 GH2-A are used to evaluate the between laboratory performance using each model. Sample GH2-01 had an HbA1c concentration of 48 mmol/mol (6.49%) and was measured by 3277 laboratories, about 90% of them in the US. The overall and manufacturer/instrument specific bias and between laboratory CV are taken from the NGSP website (18) to calculate TE and σ (Table 1, supplemental data). In Fig. 2 bias and CV are plotted in the combined BV-SM model. To visualize positive and negative bias the figure can be expanded (Fig. 1 supplemental data). The overall mean (black star) is well outside the minimum performance criterion (amber colored) of the BV model and also not within the 2-sigma limit of the SM model. The mean performance of users within each of the respective manufacturer/instrument combinations is highly variable. The majority of the HPLC instruments do, and the majority of immunochemical and POCT instruments do not meet the 2 sigma criterion of the SM model. Only one method (N) achieves the minimum performance level of the BV model. Fig. 2 also indicates the source of poor performance: high imprecision for methods A and C, high bias for methods V and W, high bias and imprecision for methods S and O. The high bias of HPLC instruments V and W is remarkable in relation to their excellent precision. The model allows historical evaluation of this phenomenon as is shown in the supplemental data Fig. 2 and Table 2. CAP data for these instruments, dating back to 2006, are evaluated. Imprecision is relatively constant but in the early years 2006-2008 (Fig. 2, blue circles) bias was low with a performance within the 2σ limits of the SM model.

Evaluation within laboratories

EQA/PT programs can be categorized  and only the highest category supplies data on bias and imprecision within laboratories (19).  To evaluate the SM and BV models within laboratories, results of such a category 1 EQA/PT program are required. CAP does not evaluate data on imprecision within labs. Data of the 2013 EQA program shared by EQA organizers in Belgium, the Netherlands, Greece and Finland are used (20). Fig. 3 shows the performances of the individual laboratories at the HbA1c concentration of 48 mmol/mol (6.5%). A substantial number of labs do not meet the the 2 sigma criterion of the SM model and many more fail even the minimum performance criterion of the BV model. In Table 3 these data are quantified. In the BV model 2, 18 and 48% of the labs meet the optimum, desirable and minimum criteria, respectively. To demonstrate the impact of chosen TAE and sigma-level, calculations for the SM model are made for TAEs ranging from 1 to 10 mmol/mol (0.09 to 0.9%) and sigmas from 2, 4, and 6. Not surprisingly, at high TAE/low sigma nearly all labs pass and at low TAE/high sigma nearly all laboratories fail. At a reasonable TAE of 5 mmol/mol (0.46%) and 2 sigma, 77% of the laboratories pass.

Discussion

Analytical/Mathematical Model

An advantage of the BV model is that once inter- and intra-individual variation are established, quality targets are fixed. But at present these parameters are not irrefutable: as stated before different values are found by different authors (10, 12, 13, 14). One publication even shows different values with different analytical methods (15). The ultimate importance of reliable data on biological variation and their impact on confidence intervals in general is described by many authors (16,17). Furthermore it is questionable whether the exact intra-individual variation can be established, because the value is in the same order of magnitude as the analytical CV of even the best methods. It is also unknown whether (narrow) biological variation data of a young, healthy, Caucasian population are applicable to other (elderly, type 1 or type 2 diabetes, well/poor controlled, ethnic) populations. A practical point of consideration is that the state of the art is quite far away from even the minimum criterion of the BV model.

A disadvantage of the SM model is that it can lead to endless discussions, without decisions, on setting an appropriate TAE and the subsequent risk to fail the set TAE (sigma). But this can also be an advantage: criteria are flexible and can be adjusted to the application (monitoring versus diagnosis), test environment (developed versus developing countries), state of the art and quality/cost issues (see considerations table 1). Another advantage is that the criteria are of the highest level of the Stockholm Hierarchy Consensus (6). This criterion is clinically and not analytically/statiscally driven and may therefore be more appealing –and thus acceptable for physicians. In developing the most appropriate quality criteria the combined model (Fig.1C) can be used but in daily practice a double target may be confusing. Balancing pros and cons, the TF- HbA1c suggests that Sigma-metrics is the model of choice.

Criteria

The goal of this paper is to define a generic model to set and evaluate quality targets. Example criteria have been used to evaluate the suitability of each of the models. The TF- HbA1c will seek collaboration with professional bodies to advocate the model and discuss/refine the most appropriate TAE and risk criteria for monitoring, diagnosis and screening of diabetes and glucose intolerance. This could be done within the framework of the Consensus Statement of HbA1c of the IDF, EASD, ADA and IFCC (7). A statement on quality targets can be added, and, like the other statements, be periodically reviewed, revised and published in diabetes related media.

Exploration of the most suitable criteria is useful as a starting point for discussion. The clinical decision limit for diagnosis (48 mmol/mol or 6.5%) is close to the upper limit of the non-diabetic reference range (42 mmol/mol or 6.0%) (9,12). Thus a small difference in HbA1c performance will have a large impact on interpretation. For example, an increase of the imprecision from 2.7 to 3.7% (1.8% to 2.5% in NGSP Units) results in a 90% increase in the prevalence of diagnosed diabetes (21). A report of the Danish Health and Medicine Authority states that biases of 0.6, 1.7, and 2.8 mmol/mol (0.06, 0.15, 0.25%) increase the number of persons diagnosed with diabetes by 28, 67 and 224%, respectively (22). When this is extended to the 2014 CAP GH2-A results it can be estimated that, if all HbA1c assays in US had been performed with instruments of the manufacturer with the highest positive bias (+4.2 mmol/mol; +0.39%), the prevalence of diabetes would be 600% higher than if all assays had been done with instruments of the manufacturer with the highest negative bias (-2.3 mmol/mol; -0.21%). This is a reason to set very tight criteria, or at least to define workable criteria to facilitate quality improvement. But one should also take into account that diagnosis of diabetes, and changes in therapy should not be based on a single HbA1c result. A second or third sample dramatically reduces the impact of analytical quality on interpretation, most evident in imprecision and to a lesser degree in bias (23). Reason, perhaps, to set more relaxed quality criteria. Some organizations have already set quality targets. From current Danish quality requirements, a maximum CV  of 2.8% (1.8%) and a maximum bias of 2.8% (1.8%) a TAE of 4.2 mmol/mol (0.39%) can be derived (21). The CAP acceptance limit of 6% is 4.3 mmol/mol (0.40%) at an HbA1c concentration of 50 mmol/mol (6.7%) (18): these criteria are close to the chosen TAE of 5 mmol/mol (0.46%) in this paper. The models in this paper can be used to set quality targets for individual labs. But also to set (desirable) quality goals between laboratories within methods (performance of users of a manufacturer) or between laboratories between methods (overall performance of all labs).

Taking into account these considerations, as well as the state of the art, TF- HbA1c suggests to initially set a TAE of 5 mmol/mol (0.46%) with a sigma of 2 for routine laboratories and a sigma of 4 for laboratories performing clinical trials where optimal accuracy and precision is a requirement. But as stated above: this is the starting point for discussion with stakeholders in the field of diabetes and of a drive towards improved quality rather than the final destination or level of quality.

Application

As demonstrated by the examples in this paper the model is suitable to estimate the quality status of HbA1c of a) a single laboratory, b) a single manufacturer, and c) a country. The model also allows comparison over time of analytical performances of manufacturers, which is helpful to select a method or to evaluate whether poor performance in a single laboratory is traceable to either poor operation in that laboratory or to the method used. For valid and fair conclusions it is important to confirm commutability and value assignment with the IFCC Reference Measurement Procedure of the samples in the EQA/PT program from which bias and imprecision data are taken to calculate TE.

References

- Task Force on Implementation of HbA1c Standardization. www.ifcc.org/executive-board-and-council/eb-task-forces/task-force-hba1c/ (Accessed August 2014).

- Westgard JO, Carey RN, Wold S. Criteria for judging precision and accuracy in method development and evaluation. Clin Chem 1974;20:825-33.

- Fraser CG, Hyltoft Petersen P, Libeer JC, Ricos C. Proposal for setting generally applicable quality goals solely based on biology. Ann Clin Biochem 1997;34:8-12.

- Bais R, Armbruster D, Jansen RTP, Klee G, Panteghini M, Passarelli J, et al. Defining acceptable limits for the metrological traceability of specific measurands. Clin Chem Lab Med 2013;51:973-979.

- Westgard JO, Westgard SA. Total Analytical Error, from concept to application. Clinical Laboratory News 2013;39 – volume 9:8-10.

- Kenny D, Fraser CG, Hyltoft Petersen P, Kallner A. Strategies to set Global Analytical Quality Specifications in Laboratory Medicine – Consensus agreement. Scand J Clin Lab Invest 1999;59:585.

- Consensus Committee. Consensus statement on the worldwide standardization of the hemoglobin A1c measurement: the American Diabetes Organization, European Association for the Study of Diabetes, International Federation of Clinical Chemistry and Laboratory Medicine, and the International Diabetes Federation. Diabetes Care 2007;30:399-400.

- Weykamp CW, Mosca A, Gillery P, Panteghini M. The analytical goals for hemoglobin A(1c) measurement in IFCC units and national Glycohemoglobin Standardization Program Units are different. Clin Chem 2011;57:1204-6.

- American Diabetes Association. Standards of medical care in diabetes – 2011. Diabetes Care 2011;34(S1): S11-61.

- Westgard QC. Minimum specifications from biological variation database. www.westgard.com/minimum-biodatabase1.htm (Accessed August 2014).

- http://www.nice.org.uk/guidance/cg87/resources/cg87-type-2-diabetes-newer-agents-a-partial-update-of-cg66-short-guideline2 (Accessed October 2014)

- American Diabetes Association. Standards of medical care in diabetes – 2010. Diabetes Care 2010; 33(S1), pS23 table 11.

- Braga F, Dolci A, Montagnana M, Pagani F, Paleari R, Guidi GC, et al. Revaluation of biological variation of glycated hemoglobin (HbA1c) using an accurately designated protocol and an assay traceable to the IFCC reference system. Clin Chim Acta 2011;412:1412-6.

- Carlsen S, Petersen PH, Skeie S, Skadberg O, Sandberg S. Within-subject biological variation of glucose and HbA(1c) in healthy persons and in type 1 diabetes patients. Clin Chem lab Med 2011;49:1501-7.

- Lenters-Westra E, Roraas T, Schindhelm RK, Slingerland RJ, Sandberg S. Biological variation of HbA1c: Consequences for Diagnosing Diabetes Mellitus. Clin Chem 2014 Sep 23. pii:clinchem.2014.227983 [Epub ahead of print] PMID:25248570.

- Aarsand AK, Roraas T, Sandberg S. Biological variation – reliable data is essential. Clin Chem Lab Med 2015;53:153-4.

- Roraas T, Petersen PH, Sandberg S. Confidence intervals and power calculations for within-person biological variation: effect of analytical imprecision, number of replicates, number of samples, and number of individuals. Clin Chem 2012;58:1306-13.

- College of American Pathologists (CAP) GH2 Survey Data. www.ngsp.org/CAPdata.asp (Accessed August 2014).

- Miller WG, Jones GR, Horowitz GL, Weykamp C. Proficiency testing/external quality assessment: current challenges and future directions. Clin Chem 2011;57:1670-80.

- www.euroreflab.com (Accessed August 2014).

- Nielsen AA, Petersen PH, Green A, Christensen C, Christensen H, Brandslund I. Changing from glucose to HbA1c for diabetes diagnosis: predictive values of one test and importance of analytical bias and imprecision. Clin Chem Lab Med 2014;52:1069-77.

- Report to Danish Health and Medicine Authority by DEKS, 21 September 2012.

- Hyltoft Petersen P, Klee GG. Influence of analytical bias and imprecision on the number of false positive results using Guideline-Driven Medical Decision Limits. Clin Chim Acta 2014;430:1-8.

Table 1. Key elements considered when developing the quality targets

| Topic | Specification |
| --- | --- |
| Anchor | Reference Method |
| A Quality Concept | Total Allowable Error Total Error derived from Bias and Imprecision |
| Analytical-Mathematical Models | Biological Variation (BV) Sigma-metrics (SM) |
| Model components a. Criteria TAE b. Risk Grading c. Evaluation of each model | Intra- and Inter-individual Variation (BV) Goals derived from the Stockholm Hierarchy (SM) Optimum-Desirable-Minimum (BV) n-sigma (SM) within laboratories (BV and SM) between laboratories (BV and SM) |
| Factors to be considered when setting criteria | Aim of the test Test Environment State of the Art Economical Affordability Practical Aspects Clinical needs |

Table 2. Mathematics for both BV and SM models

| Total Error |
| --- |
| 1. TE = ¦B¦ + 2I |
| Biological Variationa |
| 2. TAE = 1.65(f1 x CVI) + f2 CVI = 2.77% (NGSP 1.9%) CVG = 8.32% (NGSP 5.7%) |
|  | Classification | f1 | f2 | TAE at HbA1c = 50 mmol/mol (6.7%) |
|  |  |  |  | In Relative Units (%) | In Absolute Units (Concentration) |
|  |  |  |  | IFCC | NGSP | IFCC mmol/mol | NGSP % |
|  | Optimum | 0.25 | 0.125 | 2.24 | 1.53 | 1.1 | 0.10 |
|  | Desirable | 0.50 | 0.250 | 4.48 | 3.07 | 2.2 | 0.21 |
|  | Minimum | 0.75 | 0.375 | 6.72 | 4.60 | 3.4 | 0.31 |
| Sigma-metrics |
| 3. |
| Conversion Factors NGSP from IFCC |
| 4. HbA1c in Concentration Units: NGSP = 0.0915IFCC + 2.15 5. Absolute Bias in Concentrations Units: NGSP = 0.0915IFCC 6. Imprecision in CV% at 50 mmol/mol: NGSP = 0.685IFCC 7. Imprecision in CV% in NGSP units: IFCC = 1.46 NGSP |
| Symbols |
| TE = Total Error in concentration units TAE = Total Allowable Error in concentration units ¦B¦ = Absolute Bias in concentration units I = Imprecision in concentration units CVI = Intra Individual Variation in CV% CVG = Inter Individual Variation in CV% σ = Chance Failure in sigma units f1 = Factor to classify the imprecision component of TAE f2 = Factor to classify the bias component of TAE |

a CVI and CVG taken from the Ricos tables in NGSP units and converted to IFCC with equation 7

Table 3. Demonstration of the impact of different chosen criteria on the % of 342 laboratories passing for the BV and SM models.

| Biological Variation | Sigma-metrics |
| --- | --- |
| Criterion | % Labs pass | Criterion Chance Failure | TAE 10 (0.91) | TAE 7 (0.64) | TAE 6 (0.55) | TAE 5 (0.46) | TAE 4 (0.37) | TAE 3 (0.27) | TAE 2 (0.18) | TAE 1 (0.09) |
| Optimum | 2 | 6 σ | 71 | 46 | 30 | 19 | 10 | 2 | 1 | 0 |
| Desirable | 18 | 4 σ | 89 | 67 | 57 | 44 | 25 | 12 | 2 | 0 |
| Minimum | 48 | 2 σ | 100 | 95 | 89 | 77 | 60 | 38 | 16 | 1 |

Figure Captions

Fig. 1. Quality Target Models.

Imprecision in %CV  on the x-axis and bias in mmol/mol (IFCC) and % (NGSP) on the y-axis. Quality Targets are shown in colors for the Biological Variation model (A; optimum, desirable, minimum). And with lines for the Sigma-metrics model (B; 2 – 6 σ). Figure 1C combines both models.The red heart represents the performance of a laboratory with a bias of 1 mmol/mol (0.9%), an imprecision of 1.5% (1.0%) and a TE of 2.5 mmol/mol (0.23%) and meets the “minimum” performance criterion in the Biological Variation model and the “5σ” criterion in the Sigma-metrics model. Fig. 1D shows the impact of varying the set TAE and acceptable σ value on performance.

Fig. 2. Models applied to 26 manufacturer/instrument means in CAP 2014 GH2-A survey.

The mean within manufacturer interlaboratory CV is on the X-axis; the mean manufacturer absolute bias is on the Y-axis. The black star represents the overall mean of all Laboratories. The dots (laboratory instruments) and squares (POCT instruments) represent specific manufacturers with colors for analytical principles: Immunoassays (green), Ion Exchange HPLC (red), Affinity HPLC (yellow), Capillary Electrophoresis (blue) and Dry Chemistry (grey). Abbott Architect c System (A), Abbott Architect I System (B), Axis-Shield Afinion (C), Beckman AU systems (D), Beckman UniCel DxC Synchron (E), Bio-Rad D10 (F), Bio-Rad Variant II (G), Bio-Rad Variant II Turbo (H), Bio-Rad Variant Turbo 2.0 (I), Roche Cobas c311 (J), Roche Cobas c500 series (K), Roche Cobas Integra 400 (L), Roche Cobas Integra 800 (M), Sebia Capillarys 2 Flex Piercing (N), Siemens Advia Chemistry Systems (O), Siemens DCA 2000/2000+ (P), Siemens DCA Vantage (Q), Siemens Dimension ExL (R), Siemens Dimension RxL (S), Siemens Dimension Vista (T), Siemens Dimension Xpand (U), Tosoh G7 Auto HPLC (V), Tosoh G8 Auto HPLC (W), Trinity Biotech HPLC (X), Trinity Biotech Premier Hb9210 (Y), Ortho Clin Diag Vitros 5,1 FS, 4600, 5600 Chem System (Z). For more details see Table 1, supplemental data.

Fig. 3. Models applied to 342 individual laboratories.

Each dot represents an individual laboratory: HPLC (red), Immunoassay (green), Affinity Chromatography (yellow), Capillary Electrophoresis (blue). For reasons of simplicity only 2, 4 and 6σ limits for TAE = 5 mmol/mol (0.46%) of the Sigma-metrics model are shown. Yellow, grey and amber are the limits of optimum, desirable and minimum performance of the biological variation model. Data from the shared 2013 EQA program in Belgium, the Netherlands, Greece and Finland.